## Slide 1: Please delete this slide after reading through and before working on your slides
TIPS
Content provided in this lesson does not replace the i-STAT User Guides, System Manual or Cartridge Test and Information (CTI) sheets or Instructions for Use (IFU).
This power point may be modified to represent products within the i-STAT System currently in use within your facility.
Once the resource has been modified, customers will then assume responsibility for the content.*
Abbott will continue to provide and support the original resources made available on the www.globalpointofcare.abbott website.*
*Check the website regularly to ensure that you have the most up-to-date content.

## Slide 2: i-STAT 1 Analyzer
i-STAT 1 System Orientation

## Slide 3: Learning Objectives
i-stat 1 system orientation | i-STAT 1 ANALYZER
  i-STAT 1 analyzer and system components
  Anatomy of the i-STAT 1 analyzer
  Warning label locations on the analyzer
  Identify a wireless analyzer (optional)
Understand the following:

## Slide 4: System Components
i-STAT 1 Analyzer
Used to perform blood analysis testing with i-STAT cartridges, review test results, and conduct quality control (QC) testing.
Optional ability to transmit results via the i-STAT 1 Downloader/Recharger.
Lesson: Patient testing with the i-STAT 1 System | i-STAT CREA Cartridge | i-STAT 1 Learning System Classroom Training Power Points – Pfizer
For in vitro diagnostic use. For intended use and product details, visit www.globalpointofcare.abbott NOTE: Not all products are available in all regions. Check with your local representative for availability in specific markets.
©2023 Abbott. All rights reserved. i-STAT is a trademark of Abbott.
RESOURCE: i-STAT 1 System Manual, i-STAT 1 Analyzer.

## Slide 5: System Components - cont’d
i-STAT 1 Downloader/Recharger
Used to perform analyzer software updates.
Recharges the i-STAT 1 9-Volt NiMH Rechargeable Battery
Transmits results via wired network connection (Optional).
Lesson: Patient testing with the i-STAT 1 System | i-STAT CREA Cartridge | i-STAT 1 Learning System Classroom Training Power Points – Pfizer
For in vitro diagnostic use. For intended use and product details, visit www.globalpointofcare.abbott NOTE: Not all products are available in all regions. Check with your local representative for availability in specific markets.
©2023 Abbott. All rights reserved. i-STAT is a trademark of Abbott.
i-STAT Cartridges
Unit-use design
Contains sensors and reagents for testing patient samples and quality control fluids.
RESOURCE: i-STAT 1 System Manual, i-STAT Cartridge.
RESOURCE: i-STAT 1 System Manual, i-STAT 1 Downloader/Recharger.

## Slide 6: System Components - cont’d
Disposable Batteries and the i-STAT 1 Battery Carrier:
The analyzer requires two Ultralife 9-Volt lithium batteries for use with the i-STAT 1 Battery Carrier.
Lesson: Patient testing with the i-STAT 1 System | i-STAT CREA Cartridge | i-STAT 1 Learning System Classroom Training Power Points – Pfizer
For in vitro diagnostic use. For intended use and product details, visit www.globalpointofcare.abbott NOTE: Not all products are available in all regions. Check with your local representative for availability in specific markets.
©2023 Abbott. All rights reserved. i-STAT is a trademark of Abbott.
CAUTIONS1:
Use only an Ultralife 9-volt lithium battery purchased from Abbott (List Number: 06F21-26).
The Ultralife 9-volt lithium battery has a safety feature that provides protection preventing the i-STAT 1 analyzer from overheating due to component failure within the analyzer circuitry.
RESOURCE: i-STAT 1 System Manual, i-STAT 1 Analyzer.
REFERENCE: 1. i-STAT 1 System Manual, i-STAT 1 Analyzer, Power.

## Slide 7: System Components - cont’d
i-STAT 1 9-Volt NiMH Rechargeable Battery (optional):
Alternate power source that can be charged using the i-STAT 1 Downloader/Recharger.
Lesson: Patient testing with the i-STAT 1 System | i-STAT CREA Cartridge | i-STAT 1 Learning System Classroom Training Power Points – Pfizer
For in vitro diagnostic use. For intended use and product details, visit www.globalpointofcare.abbott NOTE: Not all products are available in all regions. Check with your local representative for availability in specific markets.
©2023 Abbott. All rights reserved. i-STAT is a trademark of Abbott.
CAUTIONS2:
Use only a rechargeable battery pack purchased from Abbott.
Rechargeable battery packs not recommended by Abbott Point of Care or purchased from Abbott may be susceptible to overheating and could lead to a potential fire or burn hazard.
RESOURCE: i-STAT 1 System Manual, i-STAT 1 Analyzer.
REFERENCE: 2. i-STAT 1 System Manual, i-STAT 1 Analyzer, Rechargeable Battery.

## Slide 8: System Components - cont’d
i-STAT 1 Electronic Simulator:
A quality control device for the analyzer’s cartridge signal-reading function.
Lesson: Patient testing with the i-STAT 1 System | i-STAT CREA Cartridge | i-STAT 1 Learning System Classroom Training Power Points – Pfizer
For in vitro diagnostic use. For intended use and product details, visit www.globalpointofcare.abbott NOTE: Not all products are available in all regions. Check with your local representative for availability in specific markets.
©2023 Abbott. All rights reserved. i-STAT is a trademark of Abbott.
RESOURCE: i-STAT 1 System Manual, Portable Printer.
i-STAT 1 Printer:
Portable printer used to print records from the analyzer.
RESOURCE: i-STAT 1 System Manual, Electronic Simulator.

## Slide 9: Anatomy of the analyzer
Keypad
Display SCREEN
KEYPAD KEY & FUNCTION3 | Description
SCAN | Activates the barcode scanner.
 | Used as a page key to move from one screen to the next and navigate the alphabet when the ABC key is pressed.
ABC | Used to enter alpha characters on data entry screens.
0-9 | Used to enter a number or digits on data entry screen and to select menu options and stored records.
. | Enters a decimal point or a comma separator based on customization.
 | The back-light for the display is turned on and off by pressing the “0” key for one second.
 | Used to respond to a prompt to complete an action, such as entering an operator or patient ID via the keypad.
MENU | Used to return to the previous menu and switch between the Test and Administration Menus.
 | Used to print either directly to the i-STAT 1 Printer or to the i-STAT 1 Printer attached to the i-STAT 1 Downloader/Recharger.
 | Turns the analyzer on or off.
Front of analyzer
Display screen
Keypad
RESOURCE: i-STAT 1 System Manual, i-STAT 1 Analyzer.
REFERENCE: 3. i-STAT 1 System Manual, i-STAT 1 Analyzer, Keypad.

## Slide 10: Anatomy of the analyzer - cont’d
Infrared communication window
Laser BARCODE SCAN WINDOW
Battery compartment
Top of analyzer
Battery compartment
Infrared communication window
Laser barcode scan window
RESOURCE: i-STAT 1 System Manual, i-STAT Analyzer.

## Slide 11: Anatomy of the analyzer - cont’d
WARNING LABEL
CARTRIDGE PORT
CAUTIONS4:
Propofol, Thiopental Sodium & Hydroxyurea cause interference. See manual for complete interference information.
RESOURCE: i-STAT 1 System Manual, i-STAT Analyzer.
REFERENCE: 4. i-STAT EC8+ Cartridge Instructions for Use (IFU); i-STAT CREA Cartridge Instructions for Use (IFU).
Bottom of analyzer
Warning label (i-STAT EC8+ and i-STAT CREA cartridges)
Cartridge port

## Slide 12: Anatomy of the analyzer - cont’d
Back of analyzer
Model number
Warning label (barcode scanner)
Serial number
RESOURCE: i-STAT 1 System Manual, i-STAT 1 Analyzer.
REFERENCE: 5. i-STAT 1 System Manual, i-STAT 1 Analyzer, Laser Barcode Scanner.
Model Number
WARNING LABEL
Serial Number
CAUTIONS5:
DO NOT stare into the laser aperture or the laser beam or point the laser beam at others.
Class 2 laser scanners use a low power, visible light diode. As with any bright light source, such as the sun, the user should avoid staring directly into the laser beam. Momentary exposure to a Class 2 laser is not known to be harmful.

## Slide 13: Training Complete | Congratulations!
i-STAT 1 Analyzer
You have just completed i-STAT 1 System Orientation
i-stat 1 system orientation | i-STAT 1 ANALYZER